To learn more about slings and sling management refer to Section 5 in the Safe Patient Handling and Mobility: A Toolkit for Program Development 2025 at: https://www.nvha.net/safe-patient-handling-and-mobility-toolkit/

Sling Inspection Checklist (Sample)

Consult the sling manufacturer’s instruction for correct use, maintenace and inspection of their product.

| ABC Hospital Safe Patient Handling and Mobility Program Sling Inspection Checklist |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Date: |  |  | Facility |  | Unit/Dept |  |
| Inspection completed by (name): |  |  | Job Title: |  | Company |  |
| Sling Type: |  |  | Sling Manufacturer: |  | Sling Size: |  |
| Serial number/Unique Identifier: |  |  |  |  |  |  |
| Result of Inspection |  |  |  |  |  |  |
| Sling removed from Service: | Yes No |  | Date Removed: |  | Next Inspection Due: |  |
| Comments about method of disposal; if sling sent for repair; or other remediation taken: |  |  |  |  |  |  |

Additional instructions for use of this form  can be added such as:

•	Schedule for conducting sling inspection

•	Where to access inspection forms

•	Where to return completed forms

•	Questions? Contact the SPHM Program Coordinator (name & email/phone) at:

| Component Remove from service if any of the following are found | Pass Y/N | Comments |
| --- | --- | --- |
| Sling Label: - Missing, illegible, damaged/torn, not securely attached - Color coding if used to identify sling size on label – does not match color of sling binding |  |  |
| Sling Loops, Handles: - Missing or loose stitching (Box-X stitching should be present & intact) - Tears, cuts, fraying - Clip/key attachments – cracked, bent, visibly damaged, in correct orientation - Missing loops or clip/key attachments - One or more attachment loops are not securely attached (Pull on all straps/loops in opposite direction and note if stitching becomes loose or comes apart) |  |  |
| Sling Body: Lay sling out on flat surface for inspection - Binding - not intact/missing or loose stitching/threads or fraying - Body - Cuts, tears, holes, snags, worn areas (Hold sling up to a light to check) |  |  |
| Overall: Signs of chemical or heat damage: - Discoloration, puckering, staining, e.g., from bleach - Areas that are stiff (& not part of normal design) - Melted objects or burn marks |  |  |
| Opened slings do not have documented date of first use |  |  |
| Expiration date (if present) has expired |  |  |
| Label are not clearly visible on slings in unopened packages |  |  |
| Other: |  |  |